Principle of the Test

The kit uses immunochromatography:

- The test strip contains colloidal gold–labeled equine polyclonal antibodies against snake venom.

- If venom is present in the sample, it forms an antigen–antibody complex.

- This complex is captured at the Test (T) line, producing a visible red band.

- A Control (C) line confirms proper test performance.

Test Procedure

- Add 1 drop (50 µL) of serum/plasma/whole blood into the sample window.

- Add 1 drop of test diluent.

- Read results after 2 to 20 minutes (do not interpret after this time).

Interpretation of Results

- Negative: Only the control line appears.

- Positive: Both control and test lines appear.

- Invalid: No control line appears (test must be repeated).

Performance & Storage

- Sensitivity: Detects venom concentrations as low as 10 ng/mL.

- Storage: 2–30 °C until expiry date.

- Shelf life: 24 months.

- Pack size: 50 tests per kit.

Kit Contents

Includes test devices, droppers, test diluent, silica gel, alcohol swabs, Lancet and instructions for use.